## Slide 1
FIRST EXPERIMENT UPDATE
RayStation Monte Carlo for electron skin collimators
OCTAVIUS 1500 measured profiles vs RayStation DICOM dose profiles | copper plates | central cross-plane profiles

## Slide 2
2
Fitting in the larger project
The target paper is a validation study, but this is the first structured comparison pass.
Long-term question
- Can RayStation Monte Carlo calculate dose accurately when electron skin collimators are used?
- Which regions are reliable: high dose, edge, low dose, and penumbra?
- How do material, thickness, energy, field size, extraction depth, and setup offset affect agreement?
Where we are now
- First paired OCTAVIUS/RayStation profile set is processed.
- Coverage is intentionally narrow: copper only, limited field sizes, limited offsets.
- Useful for pipeline validation and hypothesis generation; too early for broad conclusions.

## Slide 3
3
First experiment: what was compared
Matched filenames pair each measured MCC profile with a RayStation CSV dose profile.
Dimension | Current coverage | Important limitation
Energies | 6, 9, 12 MeV | Not every energy has every field size
Fields | 10x10 cm; 9 MeV also has 6x6 cm | Only two field sizes so far
Collimator | One or two copper plates | One material and two thicknesses only
Profiles | Central cross-plane at dmax | No full 3D or depth-dose validation yet
Comparisons | Raw Gy plus normalized shape/penumbra | OCTAVIUS detector response is not isolated yet
Current analysis scripts: detector_array_analysis.py -> MCC central profiles; analyze.py -> RayStation point-dose CSVs; comparison.py/raw_comparison.py -> paired comparison outputs.

## Slide 4
4
Measurement-to-model comparison pipeline
Raw Gy agreement and normalized profile shape are intentionally separated.
MCC
OCTAVIUS 1500 detector rows parsed and aligned to a common x grid
CSV
RayStation DICOM dose sampled at dmax across x values
Pair
Basename-matched MCC/CSV profiles compared point-by-point
Report
Raw Gy tables, normalized metrics, SVG overlays, batch summaries
field-center
default x alignment using 80% midpoint
raw Gy
absolute-dose comparison without rescaling
relative
shape and penumbra after normalization

## Slide 5
5
Raw Gy: CSV is high in the plateau
Across non-trial profiles, the high-dose CSV/MCC ratio averages 1.043.
1.043
mean high-dose CSV/MCC ratio
1.028-1.057
range across pairs
Chart: percent above MCC
| 6 MeV / 10x10 | 9 MeV / 10x10 | 9 MeV / 6x6 | 12 MeV / 10x10
High-dose CSV above MCC (%) | 3.158 | 4.2731 | 5.1433 | 4.4247
Slice | N | Mean ratio
6 MeV / 10x10 | 4 | 1.032
9 MeV / 10x10 | 4 | 1.043
9 MeV / 6x6 | 4 | 1.051
12 MeV / 10x10 | 6 | 1.044

## Slide 6
6
Normalized shape: plateau close, edges differ
Independent 80%-midpoint normalization makes the penumbra and field-width behavior visible.
2.58 pp
mean normalized RMS >=10%
4.40 pp
mean penumbra/shoulder RMS
-1.36 mm
mean FWHM CSV-MCC
Chart: percentage points
| 6 MeV / 10x10 | 9 MeV / 10x10 | 9 MeV / 6x6 | 12 MeV / 10x10
Profile RMS >=10% (pp) | 3.6598 | 3.0867 | 1.8004 | 2.0473
Penumbra/shoulder RMS (pp) | 6.6825 | 5.3825 | 2.5556 | 3.4602
Metric | Mean CSV-MCC | Read
FWHM | -1.36 mm | CSV field is narrower
Left 80-20 penumbra | -0.84 mm | CSV edge is sharper
Right 80-20 penumbra | -1.01 mm | CSV edge is sharper
High-dose normalized RMS | 0.75 pp | plateau shape is close

## Slide 7
7
Representative raw Gy overlays
Left: 12 MeV 10x10 one-plate on-axis. Right: 9 MeV 6x6 two-plate 2 cm off-axis. Both retain the raw Gy scale; the lower panel is CSV minus MCC.

## Slide 8
7
Normalized overlays
Left: 12 MeV 10x10 one-plate on-axis. Right: 9 MeV 6x6 two-plate 2 cm off-axis. Both retain the raw Gy scale; the lower panel is CSV minus MCC.

## Slide 9
8
Good signal; not paper-ready yet
The first experiment is most useful as a check and future motivation for direction
Good
- Matched MCC/CSV pairing and central-crossplane extraction now work end-to-end.
- High-dose normalized shape is generally close.
- Absolute offset is consistent enough to investigate systematically.
- Outputs are traceable: per-point tables, metrics, and overlay plots.
Bad
- RayStation CSV is consistently high in the high-dose region by roughly 3-5%.
- CSV profiles tend to be narrower and have sharper 80-20 penumbra.
- Coverage is sparse: no full sweep across energies, sizes, materials, and thicknesses.